- Change the DI water

  - Fill from the machine, not from a tank or squeeze bottle (can fill a clean beaker or flask and use that to fill the bottle)

- Turn on IC: button in back right (don’t turn on autosampler yet)

- Open the Chromeleon software

- In the browser window, click on instrument panels, choose the traditional one

- Check the “connect” box to connect the machine and the software

- Adjust the eluent bottle level in the software to match the actual amount in the bottle

- Open the IC machine (just pull the door)

  - We normally use the smallest sample loop – use a larger one for lower concentrations

  - If changing the sample loop, make sure to click “off” first (pump off)

  - Use tiny wrench to help loosen the sample loop

  - Don’t need to tell the machine what size loop you use, but use the same one for all runs

- Prime the pump to push air bubbles out

  - Open waste valve (bottom left valve) ½ turn

  - Click “Prime” then “ok”

  - When done, click “off”

  - Close valve

- After priming, click “Startup”

- Go to Control -> Acquisition on -> ok

  - This will draw the baseline

  - Wait until the signal is approximately 1-2 uS and stable (may take a few hours if it hasn’t been run recently)

  - If the pressure is not constant, there are air bubbles - prime the pump again

  - SRS current = 112 mA

  - Cell temp = 35oC

  - Column temp = 30oC

  - Remaining life % = On

  - Concentration = 30 mM

  - Pressure and flow rate should be set to stable

- In the meantime, prepare samples and standards (see “Preparing the Samples” section below)

- Open old SALTEx file and save as something else (with new date).

  - Prog: AS11HC anions_4mm_10min	Method: Default_anions_craft

  - Can delete rows/samples and edit sample names, but make sure to re-number the vial column if you add or delete samples

  - Make sure the last sample is the shutdown sample

- Check the method: double-click it in the browser, make sure “use recently detected retention times” is checked, double-check the standards in the amount table.

- Turn on the Autosampler (switch in back right). Operation should be set to Local (Lcl). Switch from Hold to Run before starting the batch.

- Stop data acquisition. To start, go to Batch -> Start -> Add (find your sequence in browser) - > “Start”. Watch the first sample to make sure everything is working correctly.

Preparing the Samples

- Use the numbered racks (in drawer below IC)

  - First sample starts at the black dot, and then counts up to the left (654321•)

  - The data sheet (IC run sheet) should have tray # and position

- Have two blanks at the beginning of the run, two at the end, and a 10 ppm check with one blank every 10 samples (see IC run sheet template for example of how to set up a run)

  - Prepare standards by diluting the mixed anion standard (Thermo Scientific Dionex Seven Anion Standard II, Product # 057590, 100 mL) according to the “Standard Preparation Methods” section below.

- Fill sample tubes about halfway full (below faint line on tube)

- May need to dilute some samples - make 1:100 dilutions for all PR plots (and PU plots if running October samples) by pipetting 50 uL (0.05 mL) of sample and 4.95 mL of MilliQ water into the tubes. Cap and invert twice to mix.

- Run duplicates of a few samples for QA/QC purposes.

- Put the black cap on the tubes small side up and push the cap down with the flat end of the black tool in the IC drawer (cap should be flush with top of tube)

  - Only use caps that have filters

- Load the sample racks into the left side of the autosampler, with the black dot on the right

- Can load 11 racks at a time in the autosampler

Getting the Data

- To look at the chromatogram for a sample, double-click the sample name in your run file

- Go through all the samples and standards and delete any “n.a.” peaks, make sure all peaks are delimited correctly.

  - To change the baseline, move the red line

  - To change the beginning or end of the peak, move the blue tick marks

  - To change peak properties, right click the peak (then go to “Properties”)

  - Right click to “split peak” if peaks are overlapping

  - To delete a peak, click on it in the chromatogram and hit the delete key

- May need to save and re-open to re-integrate peaks

- Copy and paste results from chloride and sulfate (separately) into Excel to save.

Helpful Hints

- Each sample takes ~12 minutes to run, so plan accordingly (the machine will run on its own, but you may need to be around to switch racks).

- After the run, check the data to make sure all measured values are within the range of the standards. If the measured concentration for a plot is above the highest standard, make the necessary dilutions and run again.

- It’s better to start in the morning so you can watch the machine for a while if needed.

- See “Quick guide to the DIONEX” methods for more info

Standard Preparation Methods

Prepare standards in the IC tubes using the volumes of standard and Milli-Q water specified in the table below. Run the standards in descending order. Be sure to use Thermo Scientific Dionex Seven Anion Standard II, Product # 057590 (100 mg/L Cl, 100 mg/L SO4).

| Concentration of Cl and SO4 | Dilution Factor | Volume of Standard | Volume of Milli-Q Water | Units |
| --- | --- | --- | --- | --- |
| 100 | 1 | 5 | 0 | mL |
| 50 | 0.5 | 2.5 | 2.5 | mL |
| 25 | 0.25 | 1.25 | 3.75 | mL |
| 10 | 0.1 | 0.5 | 4.5 | mL |
| 5 | 0.05 | 250 | 4750 | uL |
| 1 | 0.01 | 50 | 4950 | uL |
| 0.5 | 0.005 | 25 | 4975 | uL |
| 0 | 0 | 0 | 5000 | uL |